Using QHY cameras and associated software
ING seminar
Jure Skvarč
2021-04-06

---
Observing system -machines

---
Observing system startup and shutdown
- obssys
  - Selects which observing system to use: release, old, dev, test
  - does not run anything
  - can be executed in several terminals, even with different parameters
- startobssys
  - starts applications and services of the observing system
  - must be executed only in one terminal
- shutdownobssys
  - terminates observing system applications
  - always use it before startobssys

---
Observing system services
- Parameter Notice Board (taurus)
  - Allows exchange of information between applications: mimics, GUIs, TCS, camera servers
  - Everything breaks if it is not running
- Obslog (taurus)
  - Show info about exposures. Not used by other applications, but important for users
- Redis (redisserver)
  - Weather station
  - indicam, skytsar

---
Observing system services
- PostgreSQL database
  - Stores various configurations
  - TWFC and Basler run numbers
  - Started as a system service
- DEWARS
  - Server started as a system service
  - Web-based GUI
  - Used to configure and monitor the autoguiding camera (AG7)

---
QHY cameras properties
- ING instrument page: http://www.ing.iac.es/astronomy/instruments/pf-qhy/
- Model: QHY600L
- Resolution: 9576×6388 pixels
- Binning modes: 1, 2, 4
- Supports arbitrary windowing
- Readout speed (theoretical): >2 full frames/s
- Real-life readout speed: ~5 s / frame
- Shutter:
  - No mechanical shutter
  - Electronic rolling shutter (as opposed to global shutter)

---
QHY cameras properties
- Gain: 0 - 200 (setting)
  - Real gain at setting 30: ~0.4 e- / ADU
- Dark current:
  - 0˚ C: 0.014e-/p/s
  - -20˚ C: 0.003e-/p/s

---
QHY cameras properties
- Readout noise
  - Binning 1: 2.6 e-
  - Binning 2: 5.8 e-
  - Binning 4: 10.8 e-
- Note that binning is done in software, that's why increase of the readout noise as a function of binning!

---
QHY cameras QE
Maximum quantum efficiency: 87 %

---
Filters
- TWFC1: Empty, R, G, B, Sloan I
- TWFC2: u, R, G, B, H alpha, Empty, z
- Current software does not support automatic filter offset of the focus!

---
Checks before observing
Use skytsar GUI to check or change settings
- Camera temperature: -10˚ C
- Gain: 30
- Filters: Must not be “Filter #1”, “Filter #2”, ...
- Check that both cameras are shown in the menu and are selectable
- Check disk space on /obsdata using df command

---
FITS images
- Stored in daily /obsdata/whta directory
- Names: TWFC1-nnnnnnnnnn.fit and
  TWFC2-nnnnnnnnnn.fit
- Run numbers for TWFC1 and TWFC2 are independent
- Header content: camera settings, including filter name, TCS, WEAVE, SQM.

---
Unavailable features
- Fast readout
  - Problem with the QHY library
  - Overheads due to data transfer and header collection
- Streaming mode
  - Got it running at 9 Hz for 500x500 images
  - Crashes the camera
- A new library which may improve the readout times to be tested this week, but don't hold your breath

---
Error recovery
- Camera not responding:
  - Check if miniwfc1 is offline (ping miniwfc1)
    - If it doesn't work, try power-cycling, but this error is likely a show-stopper for the night
    - Has not happened yet
  - INDI server problem
    - Use shut account to restart miniwfc1
    - Restarting miniwfc1 will automatically restart the server
    - You need to restore filter names using qhyreset command
    - Check cooling
    - Happened once during the night in last 6 months

---
Error recovery
- Camera crash
  - indicam list only shows one camera (or 0)
  - needs camera power-cycling
  - only happened during day time when testing the QHY library

---
Error recovery
- Cameras not at -10˚ C:
  - Using skytsar turn the cooling off and on and check that the cooling is activated (power goes to high values until the temperature is reached). You may need to repeat few times.
  - Do not forget to enter the set point
- Filter names not present
  - Means that the server has been restarted
  - Use qhyreset to restore filter names

---
Command line tools
- qhyreset
- indicam
- autofocusQHY (was autofocusWEAVE)
- spiral_search (will not be shown)
- tel_offset (will not be shown)
  All command line tools have option -h (the same as --help), which will show basic usage

---
qhyreset
- Does not really reset the cameras
- To be used after indi server restart
- Initializes some items in the server
- Configures filter names
- Sets the cooling set point (does not work well for TWFC1)

---
indicam
- List cameras:
  - indicam list
- Make exposures (run, dark, bias):
  - indicam run TWFC1 10
- Change filters:
  - indicam filter TFWC1 -f “Empty 1”
- Allows multiple exposures:
  - indicam -n 10 dark TWFC1 10

---
indicam
- Set binning and/or windowing
  - indicam -b4 -w “[4000:5000,3000:4000]” dark TWFC1 10
  - default binning is 2!
  - Window coordinates are in binning 1 units!
  - Default size is full frame
- Caution: Binning and readout window must be set for each command, except when using defaults! Previous settings are always overwritten.
- Error in on-line documentation: http://www.ing.iac.es/astronomy/instruments/pf-qhy/PF-QHY-command-summary.html

---
autofocusWEAVE → autofocusQHY
- https://ingconfluence.ing.iac.es:8444/confluence/display/WEAVEDEV/Procedure+for+autofocusWEAVE+focus+run
- Above link is a bit out of date, will be updated
- New:
  - Option --noprompt. Will proceed with autofocusing without confirmation
  - Option -p (--plot). Will use fitPanel to show the measurements and possibly refit the curve

---
autofocusQHY
- Example command:
  - autofocusQHY --expTime 7.0 -x 705 -y 740 -p -s TWFC1 0.0 1.0
- Exposure time 7 s (--expTime 7.0)
- Star coordinates x=705, y=740 (-x 705 -y 740)
- Plot the measuremens after fitting (-p)
- Save images (-s)
- Use camera TWFC1
- Focus range 0.0 to 1.0 mm
- Number of positions: 9 (default value, for other values use option -n)

---
autofocusQHY
- Simulation mode (option --simulation) allows processing of existing images
- Copy focus run image to a directory just for these images
- It reads focus positions from the header
- Camera name and first and last focus positions are required in the command line, but ignored
- See demo

---
GUIs
- skytsar
- plotnfit

---
skytsar
- A tool for image acquisition, display and analysis
- Acquires DAS, QHY and Basler images
- Automatic detection of available cameras
- Grab and Save modes
- Readout window selection
- Binning selection

---
skytsar
- TV mode
- Colour mode (Basler only)
- Gain control (where camera allows it)
- Filter control (QHY only)
- Expose (run), dark, bias modes of integration (reflected in FITS headers)

---
skytsar
- Zoom
- Grey (or colour channel) histogram
- Automatic or manual image level control

---
skytsar
- Quick analysis options
  - 2D profile:
    - Centroid coordinates
    - Minor and major axis, rotation
    - Total intensity
    - 80 % energy diameter, FWHM
  - 1D profile:
    - Amplitude
    - Background
    - FWHM

---
skytsar
- Quick analysis options (continued)
  - Display of values
    - 9x9 grid
  - Statistics
    - Minimum and maximum value
    - Mean
    - Standard deviation
    - Skewness, kurtosis

---
skytsar
- Quick analysis options (continued)
  - Quick offset to the target position
  - Requires correct camera rotation and pixel scale parameters in the configuration file.
  - Useful mostly for pointing test

---
skytsar
- Telescope tab
  - Basic telescope control (goto)
  - Offset
  - Guide star selection

---
skytsar
- Measurement tab
  - Contains measured quantities
  - Measurements can be selectively deleted
  - Measurements can be saved

---
skytsar
- Autofocus tab
  - Do not use yet!
  - Front end for autofocusing script
  - Should provide graphical interface for input parameters and show star isophotes and profiles
  - Problems with GUI responsiveness
  - So, do not use yet

---
skytsar
- Parameters tab
  - Avoid changing
  - Cannot overwrite default configuration file (for safety)
  - It is possible to change parameters, which are then valid until the end of session or until they are changed again

---
skytsar
- Problems
  - Somewhat slow
  - Some features not finished or well-tested
  - User experience issues
  - Lacks some useful features
  - Needs restart when INDI server is restarted
  - Sometimes ignores errors quietly

---
skytsar
- Future upgrades:
  - Autofocus
  - Autoguider
  - FPI camera interface
  - Overlay management
  - More options for image levels scaling
  - Image server access
  - Show FITS header
  - Handle multiple images
  - Web browser, email client, MP3 player...

---
plotnfit
- Used for quick plotting'n'fitting.
- Has fixed selection of models
- For some models computes certain derived quantities
- Allows interactive deselection/selection of points
- Not meant to be a replacement of usual tools
- Demo