Instrument performance sheet

	Instrument
	Wyffos

	Dates scheduled
	2  -13 january with AF2


	Preparation
	Instrument was tested before instrument change day. Discovered that one of the panamech couplings on the lamp slide had snapped.




	Day
	30 december

	Faults:
	

	Fault title:
	

	Time lost:
	This is the second time that this has happened now.

	Description
	The lamps in the calibration unit are fitted on a slide that is moved by a motor. An extra brake is fitted to hold the mechanism in place

We investigated and found that the motors and brakes are operated from two different crates. The brake releases when  it is powered up, therefor the crate has te be on! This means if the brake crate is switched off and you move the motor, the coupling will snap. 



	Solution
	Fitted a new panamech coupling


	Summary


	As long as both crates are switched on nothing will go wrong. But if we have a bad powercable or a blown fuse than things will go wrong.



	Recommendations


	Is it possible to interlock the motor movement with the crate for the brakes? 

Maybe we can power the motorcrate through the brake crate: so if the brake crate is not powered on  than the motor can’t move either.

Opsmeeting agreed that AGB should look into this.

	Needed spares


	Spare Panamech coupling for wyffos calibration unit is ordered.



	
	


Last updated by mvdh 15 jan 2003

