From rgmr@ing.iac.esWed Feb 12 16:46:25 1997

Date: Wed, 12 Feb 1997 16:11:24 +0000 (GMT)

From: Rene Rutten <rgmr@ing.iac.es>

To: Peter Moore <pcm@ing.iac.es>

Subject: CCD TV system

Hi Peter,

here's draft version of specs I started to draw up for TV CCD systems.

It never got much further than that. I've also been discussion these

issues a bit with people at RGO, where they are looking into a similar

developement for Gemini. We may want to piggy-back on that.

R.

-------------------------------------------------------------------

Dear all,

Stuart and I have looked at the specifications for CCD-based

TV systems which will replace the TV systems that are currently

in use at the ING. In order to write these requirements we

first assessed the performance of our current system, taking

the WHT ISIS slit-view TV as a working base line.

The new CCD TV system as described below will have a much

broader functionality than 'just' viewing the sky; the

additional possibilities will be of great benefit to observers

and allow persuing scientific objectives in their own right, or

supplement data taken with the main instruments.

This is only a first draft; comments are appreciated !

Stuart and Rene'

==============================================

Specifications for future CCD-based TV systems

==============================================

area of application of the TV systems

-------------------------------------

The future CCD-TV systems will be designed to replace the

current integrating tube TV systems. The TV systems are in

use at the WHT Cassegrain focus for ISIS slit viewing and

direct sky viewing, in the Nasmyth foci for UES slit viewing

and at GHRIL for sky viewing and autoguiding. At the INT TVs

are used for IDS slit viewing and in the guider telescope.

At the JKT TVs are used for sky viewing with CCD imaging, and

for slit and direct sky viewing with RBS.

Apart from the their main functionality, the TV systems are

also used for seeing measurements, collecting engineering

data, and occasionally for autoguiding,

This functionality, and more, has to be incorporated into the

new CCD-based TV systems.

Performance of the current TV systems

-------------------------------------

The optical quality of the existing TV systems is rather

poor. The tubes suffer from severe distortions which renders

the images unsuitable for accurate measurement of spatial

distances for scientific and engineering purposes. The cosmetic

quality of the image is also poor due to periodic striping,

electronic interference and a non-uniform background. The TV

systems also suffer from a poor dynamic range of only 6 bits,

as measured on actual digitised images, their quantum efficiency

is poor compared to CCD systems, and there is the constant risk

of damage due to excessive light levels.

Apart from these fundamental problems, our systems suffer from

limited functionality due to the design of their control system

which renders them extremely inflexibility (e.g. try to grab,

store and transfer a digitized image !). ING supports different

systems on the three telescopes, which adds to the support

overhead.

In spite of all these negative points, the existing TV systems

do provide a working system which is fairly robust, pleasant

and easy to work with, at least for the standard applications.

Performance requirements for the CCD-TV systems

-----------------------------------------------

sky-projected pixel size - better than 0.2 arcsec/pixel

field-of-view            - depending on focal station

chip size                - minimum 500 x 500

dynamic range            - 10 bits, minimum of 8 bits

read noise               - better than 10 electrons

response time            - variable at least from 0.2 to 100 seconds

gain                     - variable from 1 to some large number

to allow use of full well capacity with

dynamic range

additional requirements are:

1. real-time display of images

2. photometry on at least three user-defined sub-windows

3. automatic flat fielding (optional)

4. storage as FITS files (optional)

5. good blue response (optional)

6. on-line bias subtraction, flat fielding and masking of

cosmetic defects (optional)

7. readout of any sub-window

8. on-chip binning (optionally)

9. built-in centroiding on point source

10. autoguiding on point source (integration with TCS)

11. logging of centroiding errors for engineering purposes

12. automatic positioning of an object on pre-defined CCD pixel

by offsetting the telescope (integration with TCS)

13. color and range of ND filters to allow bright stars

14. safe to high light levels

Comparison of the current performance of the TV systems against

CCD images taken under identical sky conditions show that a

CCD-TV system should double or even tripple the S/N of stellar

images. Limiting magnitudes should improve substantially (by at

least one, possibly two magnitudes, depending on sky brightness).

Images from a CCD-TV system will become an important tool in

their own right, particularly if they can be pre-processed and

archived. For instance, photometry will allow continuous

measurement of important parameters such as seeing, sky

transparency, and slit losses. In particular the latter will be

of great benefit to all spectrophotometric work.

Hardware environment

--------------------

Software environment

--------------------

Rene' Rutten

Stuart Barker

1 April 1996