###############################################################

#     this file records what has been done during  NAOMI      #

#     commissioning at WHT (La Palma )                           #

##############      X.Gao 10~19/08/2000   ########################

CCD-SDSU controller

10/08/2000

1). checked the video cable, found out that Slave cable in 50way connector end

has one CCD output wire's plaster shell (cut) out by the fixing

clamps.  Ask electronics workshop fixed, put insulation tape around. Also tidy

up others.

Software

see READMEfirst.doc

10/08/0000

1). tested iraf, can display image, imexam, works

2). tested QI_server, works with VMEdummy_star(_epics)

11/08/0000

3). tested gmake for *.vws  works, change startup_cam from simulation mode to

real mode ( can not re-write, so, named another startup_cam_epics_xg.vws)

4). tested gmake for *.m4  works, add VME*.m4, works

5). copy scripts of xg* from atc /sw4/naomi/naomiWfs/scripts/xg to aocontrol1

/naomi_wht_xg/scripts_atc

6). put /sw4/naomi/naomiWfs/dsp_code/xg/DOWNLOAD_MASTER_DUMMY_STAR.lod   and

/sw4/naomi/naomiWfs/dsp_code/xg/finalversion/vmeapl4_dwld.lod

into /sotfware/naomi/naomi-0.2/naomiCam/dspsrc/vme and  into

/sotfware/naomi/naomi-0.2/naomiCam/bin/asm56000/timing/

/sotfware/naomi/naomi-0.2/naomiCam/bin/asm56000/vme/

12/08/0000

7). asked Andy.V to change permission for M4, delete ABT in (close1.m4, close2.m4) to Timing

as VME code already done it when it receives ABT. gmake

8). asked Andy.V to change permission for vws, change startup_cam (sdsuLib from simulation

mode to real mode). gamke

9). asked Andy V to copy cameraTop.db to /software/naomi/naomi-0.2/naomiCam/./data/ and tested

startup_cam, things are OK tested with VMEdummy_star_epics

10).tested with SDSU controller on <readout_cam1/2 works

11).measured the noise with INGRID SDSU switched off, slaver seems good, master slight high

/naomi_image/testimage/ master1.fits (SET=10, HIH), master5ms.fits (SET=200,HIH), slaver5ms

(SET=200,HIH), the cooler set to 8 degree. but the iraf would not let me change paramter

13/08/0000

12).the INGRID is switched on with side lid on ( hot ). the INGRID  GND was not connected yesterday, now it is.

13).fixed the problem with iraf, it is working fine. take noise measurement again (sassume gain=0.6)

/naomi_image/testimage/ master5msHihIngird.fits (5.57, 4.67, 5.27, 8.71)

master5msSlwIngird.fits (4.81, 3.94, 3.99, 5.3 )

slaver5msHihIngird.fits (4.69, 4.12, 4.6, 4.89 )

slaver5msSlwIngird.fits ( 4.15, 3.43, 3.64, 3.79)

only one channel of the Master is bit higher 8.71, the rest is ok.

14/08/0000

14). tidy up the drawing for cable and GND, now it is called  aocontrol1 /home/xg/naomi_wht_xg/ naomicable.xls,

also copy Naomicable.xls into

cp Naomicable.xls /software/naomi/naomi-0.2/docs/software/naomiCam/dsp_code/vme_board/

remove gnding1, gnding2 VSD file, remove naomicable2.xls in here too

copy alba /act_naomi/naomi-0.1/docs/camera/ naomi_ICD_parttwo.doc and naomi_ICD_partthree.doc to

aocontrol1 /home/xg/naomi_wht_xg/, also to

/software/naomi/naomi-0.2/docs/software/naomiCam/dsp_code/vme_board/

15). check the GND between CCD header and Controller, slaver==header =0.6 Ohms, Master==header =23 Ohms

put an exteral Earth wire, the Ohm all ==0.3. the Ingrid controller now is put on top of Naomi

and the side lip is open, with Ingrid switch on, Master quadrant 1 noise =18ADU, and =15ADU

when Ingrid is switched off. the salver still OK. Wait for another time to test ( change cable ?)

15/08/0000

16). the Ingrid controller is off the Naomi controller, but Ingrid has add to the cooling circulation

made a change to

Master controller =====slave cable==== master header

slave controller  =====master cable==== slave header

the noise test (HIH) results are master ( 6.4, 4.56, 6.72,9.48), slave (4.62 max) assume gain=0.6e/ADU

17). measure the dark current ( assume gain=0.6e/ADU)

master:  swpedcam12s.fits, swpedcam110s.fits,  15 e/pixel/second

slave :  swpedcam22s.fits, swpedcam210s.fits,  15.6 e/pixel/second

16/08/0000

18).take gain measure, and ask Derek for naomigain ( I could not run at here)

//////////////////////////////////////////////////////////////////////////////

Check (5:35,5:35] in each set you send me and got the following results:-

slave gain for that quadrant = 0.53 e/ADU

master gain = 0.55 e/ADU

///////////////////////////////////////////////////////////////////////////////

take noise measurement again

a). change back with Ingrid controller working

/home/xg/naomi_image/testimage/

Mater controller===master cable=====master header,  mastertry0.fits (HIH,SET=200)

Slave controller===slave  cable=====slaver header,  slavetry0.fits (HIH,SET=200)

noise is same as before

b). change back with Ingrid controller off, not see much difference

Mater controller===master cable=====master header,  mastertry1.fits (HIH,SET=200)

Slave controller===slave  cable=====slaver header,  slavetry1.fits (HIH,SET=200)

switch on Ingrid controller again

c). swap clock cable for master and slave, the video cable kept unchanged,not make much difference

Mater controller===slver  clk cable=====master header,

Slave controller===master clk cable=====slaver header,

d). swap controller, get much better noise

Mater controller===slave  cable=====slaver header,  	msCtrlSlhdHih.fits (HIH,SET=200),

msCtrlSlhdSlw.fits (SLW,SET=200)

Slave controller===master cable=====master header 	slCtrlMshdHih.fits (HIH,SET=200),

slCtrlMshdSlw.fits (SLW,SET=200)

e). change Slaver controller Timing board  setup into master setup ( Clk, EPROM, sync )

change Master controller Timing board  setup into slave setup ( Clk, EPROM, sync )

test noise again, get similar to d). Now, the configuration (re-labeled)

Mater controller===master cable=====master header,  	newmasterHih.fits (HIH,SET=200),

newmasterSlw.fits (SLW,SET=200)

Slave controller===slave cable=====slaver header 	newslaveHih.fits (HIH,SET=200),

newslaveSlw.fits (SLW,SET=200)

The NOISE results ( take master/slave gain=0.55e/ADU)

imstat 				[6:36,44:74]  [53:83,44:74]  [53:83,6:36]  [6:36,6:36]

newmasterHih.fits	4.79		4.16		4.27		5.45

newmasterSlw.fits	4.08		3.07		3.23		3.71

newslaveHih.fits	4.77		4.15		5.04		4.79

newslaveSlw.fits	4.04		3.48		3.60		4.82

msCtrlSlhdHih.fits	4.53		3.99		4.76		4.93

msCtrlSlhdSlw.fits	3.84		3.15		3.67		4.77

slCtrlMshdHih.fits	4.48		3.75		4.06		5.21

slCtrlMshdSlw.fits	4.08		2.99		3.29		3.69

mastertry0.fits		4.91		4.25		4.71		7.52

slavetry0.fits		4.11		3.83		4.07		3.91

OK, the noise at high speed under new configuration is near as good as at ATC test. fine turning

needs much more access time of WFS camera, which is not possible at the commissioning stage. Maybe

some test when mechanism control is running.

17/08/0000

19). write Camera operation/test proceducret, tidy-up the docs, get docs printed out for tomorrow. wait

for test when motor is going to move. Had a go with motor runnning. only very occasion see a sudden

noise on master, but later repeat test did not see. otherwise, the noise is not changed much. the slave

seems bit noiser while master become better than before, could be caused by someting be rearranged

on optic bench. need more time to paly with.

18/08/0000

20). handle over meeting, demonstration, modify READMEfirst.doc and Naomi_wfs_camera_op_testing.doc

test startup_cam, with close restart, it works well. tidy-up *.m4 and back-up

cp /software/naomi/naomi-0.2/naomiCam/dspsrc/scripts/*.m4  into /home/xg/naomi_wht_xg/script_wht

21). noise calculation for last night.

mastertmttest2.fits		4.16		3.92		4.16		4.46

slavemttest2.fits		5.68		3.99		4.90		6.92

19/08/0000

22).  the camera is used by others most of the time , only able to try mode2, 3 to see the image, did notice

that in mode 2, the image updateing rate in SAOtng window is not as fast as in mode1, do n't know why,

it could be the data is so fast that Vxwork could not couple, so ,drop some.

Issuers need attention

1).  the Peltier cooler power supply fuse has been burnt the second time, this needs to be investigated.

as fas as I kwon, this should not happen, it is either the fuse or the large current. the CCD

can take maximum 4A.  Need to check with either Derek or Brenda  for the final circuit.

2). the loose Lemon socket on Master header for Peltier power needs to be fixed, if the loose nut inside the

header touch the wire, it could be the cause of burnt of the fuse.

3). the scripts have to be checked for final release thought I had tidyed up, there needs a high level camera

start script.   readouts with epics have been tested with QI_server, start then stop, start  again or start

another cam ,it works well. It should be no problem with C40, but has not been tested.

Suggestion

1). C40 system should check whether there is frame missing and log message at least. Chatted

with Richard, he says, at present, C40 code is not checking, but he will add.

the c40 system should know if the data sent from CCD controller is not shifted at

pixel level, this can be tested by deliverly shifting image in Timing code.

2).  sdsu_Rs422 converter card should have two 14bit ( DIP switch )

A. for bright star, using high 14 bits

B. for normal operation using low 14 bits

3). sdsu_Rs422 converter card should have a DIP switch to connect RMT_RST to FIFORD or

dis-connect RMT_RST to FIFORD, as by disconnecting RMT_RST to FIFORD, the VME I/F DSP

should be able to send data to C40 directly without any CCD and controller. It has been

implemented in VME I/F DSP code, and tested with QI_server. But not tested with

C40.  With this facility, C40 system can do a simulation test without running camera

and can test its centroid algorithm, can check if it can detect shift at pixel level by

setting the dummy_star in such a way that the spot position is shifted deliverly.