From sgr@ing.iac.es Fri Nov 12 15:25:29 1999

Date: Wed, 03 Nov 1999 17:39:45 +0000

From: "Simon G. Rees" <sgr@ing.iac.es>

To: Gordon Talbot <rgt@ing.iac.es>, Kevin Dee <kmd@ing.iac.es>,

Chris Packham <cp@ing.iac.es>, Paul Jolley <jolley@ing.iac.es>,

Matthieu Bec <mdcb@ing.iac.es>, Peter Moore <pcm@ing.iac.es>

Cc: Frank Gribbin <fjg@ing.iac.es>,

IMPB Software Project Correspondence <impbsoft@ing.iac.es>

Subject: Thermal Cycle 1: Report on INGRID Mechanism Tests

Dear all

Here's the results of the mechanism tests Paul and I

have performed, both before cooling INGRID down, and

afterwards.

I won't bore you with the fine details, I'll just

present the major conclusions. If you want to look

at the raw test data it's available in the Software

Project File Area (entrance to 4th floor on right hand

side). Feel free to take a copy, but please leave the

masters.

Overall summary: within the constraints of testing

that was conducted _visually_ it looks like we can

get a system that's reliable when running cold.

Worst case wheel move times take longer than we

would like.

Before Cooling

--------------

During this phase of the testing we focussed mainly

on the Pupil Stop Wheel as this has the most stringent

repeatability specification (and being at the front

of the window is easiest to see !).

Within the limitations of what we could determine by

_eyesight_:

1.  The default settings Matthieu used for exercising

the test jig also appeared to work reliably and

repeatably for INGRID 'the instrument'.

These settings specify a max velocity of 4.22

(don't know what the units are) and acceleration

time of 0.9 seconds.

[Note: the sample size was limited to about 15

moves - we could really do with some 'soak'

testing. I'll try to fit this in before warming

up again - SGR].

2.  With a max velocity of 15 (ie 3 times the default)

the  pupil stop wheel was starting to _occasionally_

fail (lock up for a bit, then free itself)

3.  At a velocity of 35, the pupil stop wheel was

_consistently_ failing.

4.  We couldn't seem to improve the reliability/settling

times by changing the acceleration rates.

5.  A worst case move (index position 11 -> 1) took 86s.

This is slower than we would like.

6.  Other mechanism wheels seemed to work ok, but were

not tested thoroughly.

We didn't exercise the Pupil Imager as it wasn't

fitted, nor the Focus Drive (as Paul didn't want to

for some reason which I forget).

After Cooling

-------------

1.  The default pupil stop motor settings all appeared to

work reliably at cold temperatures too. A similar

number of moves were sampled.

2.  The velocity at which the pupil stop wheel appears

to start failing is reduced (VEL = 10). Thus our

'reliability margin' at cold temperatures may well

be reduced.

3.  Could get the system to work reliably at VEL = 6.

At this velocity a worst case move took 63s.

4.  We tested all the other mechanism wheels at the default

motor settings. All were ok. When running them at higher

velocities they didn't seem any more reliable than

the Pupil Stop Wheel (although the testing in this

area was very much reduced).

Any queries / comments / ideas - please get back to me

or Paul.

Best Regards

Simon