Isaac

Newton

Group

Red

Imaging

Detector

Release notes for SDSU software Version 2.1

Peter Moore 16th August 1999..

Version 2.1 of the INGRID SDSU Timing board software is built upon Version 2.0 and incorporates minor changes to fix bugs, remove artefacts, and improve the software behavioural stability. It also serves to familiarise the author with the detector controller software. Please read SDSU INTERFACE CONTROL DOCUMENT VERSION 2.1 in conjunction with these notes.

|  | Ver | Bug Report | Current Status in V2.1 |
| --- | --- | --- | --- |
| 1 | 2.0 | MRA Command accepts ZERO as valid input for number of readouts. Should either reject this with error or behave normally returning an image of size zero bytes. Currently it streams data in image mode. | MRA Command now accepts zero as valid number of reads and returns image of zero pixels. This can be used gainfully to test scripts without actually acquiring data. |
| 2 | ICD | ICD provides no way of setting required detector servo temperature. | See SDT command in ICD v2.1 |
| 3 | ICD | How are version numbers encoded in 24 bit memory locations ? | As compact 3 character ASCII format. Version number ‘000’ is reserved to indicate that no application program is present. |
| 4 | 2.0 | Software returns strange version numbers. | Version 2.1 of software now consistently returns string ‘2.1’ as version for boot and application code at locations $P:6 & $P:7 respectively for utility and timing code. |
| 5. | ICD | How is the temperature encoded in the 24 bit memory locations ? | Now returns temperature in units kelvin. See TEM command in ICD 2.1 |
| 6. | ICD | Specification for TEM command gives no way of checking whether command was accepted. Unlike with RDM and CHK commands which are resident in the boot code, this is an issue. | TEM command will either return an integer in the range of 0 – 333 Kelvin (hex 000000 – 00014D) or ascii code ‘ERR’. |
| 7 | 2.0 | Attempt to read minimum exposure time from X:30 gives HED error. | ICD gave incorrect location to see minimum exposure time. See list of current locations below for correct location map. |
| 8 | ICD | ICD doesn’t define mapping between physical location of pixels and order in which they are transmitted back to host. | No change – awaiting characterization efforts to arrive at optimum setting. The pixel order may change. |
| 9 | 2.0 | ABR command doesn’t abort until middle of MNDR sequence. The preferred behaviour is that ABR will be processed at each end of array read. | ABR now asyncronously aborts a readout during the integration stage of any exposure sequence with correct reply sent (DON). This implies that the das system will be able to predict the exact amount of data that arrives (i.e. one complete MNDR sequence) to facilitate host FO switching (data / msg). The ABR command cannot abort readouts during an MNDR sequence due to limitations of the SDSU architecture. |

Figure 1 Bug Fix status for Version 2.1.

Changes to Behavioural model.

- The Command sequence used to generate test data has been modified. The command DAT is now sent with one mode select variable between the ranges of 0 and 3 that corresponds to real image data, test generator 1 or test generator 2 respectively. In addition a 3rd mode (3) is incorporated that allows ‘read up ramp’ mode. This mode takes n integrations (where n is the variable n of the MRA command) of exp time p each.  The variable p is the exposure time set by the ‘SET’ command. This sends one frame of data back after reset and then each p milliseconds for a total of n frames. The Default power on (and after CON command sent) is zero, real image data. Once the mode has been set it will remain in effect until cancelled by setting another mode value  The readout mode currently selected can be read back from the timing application at address $X:3A

- The MRA command now accepts as valid multiple readout values integers between 0 and 16,777,216 (2^24).

- A new command at the boot level is incorporated to allow simple testing of communication integrity to/from the sdsu controller. By issuing the command ‘NOP’ (no operation) the sdsu will respond with ‘DON’ (done). This command cannot unfortunately be used during a MNDR sequence as it compromises the timing integrity and competes for the fiber comms link (limitation of the SDSU controller), however, it will be queued and processed after an MNDR sequence finalizes.

- Checksum values for code are now as follows:-

Timing Boot			B46057

Timing Application		61A2FA

Utility Boot			EE8C33

Utility Application 		323D77

Noted inconsistencies from version 2.0 for evaluation and ratification.

- Exposure timer is set running after completion of post reset  MNDR sequence. This means that the effective integration time is the time for one MNDR sequence (Min exp time @ X:38) plus whatever integration value is set with the SET command. The minimum exposure time from the point of view of the host is therefore zero millisecs which will result in an integration equal to the minimum integration time of 0.847 secs. This inconsistency must be dealt with to provide accurate (0.1%) timing information to the user.

- After boot or reset to the sdsu, all memory locations read back to the host will be at a value of 0xFFFFFF. This may confuse the host system if not aware that the application code is not available. This could happen by a spontaneous boot from the watchdog, etc.

.

| Board | Address | Available | Type | Description |
| --- | --- | --- | --- | --- |
| Timing | $P:6 | Boot | ASCII | Timing boot eeprom sw version |
| Timing | $P:7 | Appl | ASCII | Timing application code sw version |
| Timing | $X:0 | Appl | INT | Timing Application Mode Status Word (1) |
| Timing | $X:1 | Appl | INT | Target Integration time |
| Timing | $X:2 | Appl | INT | Elapsed Integration time |
| Timing | $X:30 | Appl | INT | Number of MNDR Reads in current exposure |
| Timing | $X:31 | Appl | INT | Delay between reset and first read. |
| Timing | $X:35 | Appl | INT | Number of reset pulses per pixel column |
| Timing | $X:36 | Appl | INT | Number of reset cycles per reset |
| Timing | $X:37 | Appl | INT | Time in ns for each pixel read |
| Timing | $X:38 | Appl | INT | Time in ms for full frame read (i.e min exp time) |
| Timing | $X:39 | Appl | INT | Number of preconditioning cycles per read |
| Timing | $X:3A | Appl | INT | Current readout mode (0 = real, 1 & 2 test modes) |
| Utility | $P:6 | Boot | ASCII | Utility boot eeprom sw version |
| Utility | $P:7 | Appl | ASCII | Utility application code sw version |

Figure 2 LIst of active addresses within the SDSU enviroment.

Notes:

1. Status word defined as:

| Bit # | Significance |
| --- | --- |
| 0 | Application executive loop active |
| 1 | Continuous reset to detector active |
| 2 | Continuous video streaming mode active |
| 3 | Clock test mode active |
| 4 | Readout / integration in progress |