INGRID ARRAY and CONTROLLER WIRING

The electrical connections required to operate the HAWAII array inside the INGRID cryostat using a SDSU controller are described in this note. The grounding and earth configurations are also described.

The connection system uses a star point configuration for grounding purposes, as far as practically possible, to reduce the problems associated with ground loops. This star point is in the SDSU controller itself. Most of the signal lines are shielded but with the shield only connected to the ground at one end only. Again this ensures no ground loops are formed. This configuration acts best to reduce low frequency pickup – mainly 50 Hz hum and its associated harmonics. It works less well for high frequencies where shielding ought to be connected at both ends.

This note shows the earthing and grounding configuration in a few very level diagrams. These diagrams should be used in conjunction with the tables which give all the details to the cabling.

This note only describes those cables associated with driving the HAWAII array. It does not give the details required to operate the moving mechanisms etc.

The connections and configurations can be broken down to the following sub sections.

- SDSU Controller internal and front panel connections

- SDSU Controller front panel to INGRID Preamp Box connections

- SDSU Controller clock cable connections

- INGRID to HAWAII fanout board connections.

- INGRID safety earth connections.

INGRID SDSU CONTROLLER INTERNAL GROUND CONNECTIONS

# PSU

SDSU FRONT PANEL

VIDEO 1 BOARD

VIDEO 2 BOARD

UTILITY BOARD

CLOCK BOARD

SDSU FRONT PANEL INTERNAL CONNECTIONS

( connections from front panel connectors to boards in controller)

| SIGNAL NAME | FRONT PANEL 50 way D | FRONT PANEL 26 Way |  | VP1 BOARD 9 way D | VP1 BOARD 15 way D | VP2 BOARD 9 way D | VP2 BOARD 15 way D | CLK BOARD 37 way D | WIRE TYPE |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| OUT1 | 2 |  |  | 4 |  |  |  |  | COAX |
| OUT1 SHIELD | N/C |  |  | 5 |  |  |  |  | COAX SHIELD |
| OUT2 | 3 |  |  | 2 |  |  |  |  | COAX |
| OUT2 SHIELD | N/C |  |  | 7 |  |  |  |  | COAX SHIELD |
| OUT3 | 4 |  |  |  |  | 4 |  |  | COAX |
| OUT3 SHIELD | N/C |  |  |  |  | 5 |  |  | COAX SHIELD |
| OUT4 | 5 |  |  |  |  | 2 |  |  | COAX |
| OUT4 SHIELD | N/C |  |  |  |  | 7 |  |  | COAX SHIELD |
| AGND | 19 |  |  |  | 8 |  |  |  | 22AWG |
| AGND | 20 |  |  | 8 |  |  |  |  | 22AWG |
| AGND | 21 |  |  |  |  |  | 8 |  | 22AWG |
| AGND | 22,35 |  |  |  |  | 8 |  |  | 22AWG |
| AGND | 35 | FLYING |  | LEAD | connected | to | CHASSIS | only | 22AWG |
| +15V | 1 |  |  |  | 9 |  |  |  | 22AWG |
| BIASGATE | 50 |  |  |  | 4 |  |  |  | 26AWG |
| HIGH3 | 7 |  |  |  |  |  | 1 |  | 26AWG |
| OFFSET1 | 18 |  |  |  | 1 |  |  |  | 26AWG |
| HIGH2 | 23 |  |  |  | 6 |  |  |  | 26AWG |
| HIGH4 | 24 |  |  |  |  |  | 2 |  | 26AWG |
| -15V | 34 |  |  |  | 10 |  |  |  | 22AWG |
| OFFSET2 | 36 |  |  |  |  |  | 3 |  | 26AWG |
| OFFSET3 | 37 |  |  |  |  |  | 4 |  | 26AWG |
| OFFSET4 | 38 |  |  |  |  |  | 5 |  | 26AWG |
| HIGH1 | 39 |  |  |  | 3 |  |  |  | 26AWG |
| BIASPWR | 17 |  |  |  |  |  | 6 |  | 26AWG |
| VRESET | 41 |  |  |  | 2 |  |  |  | 26AWG |
| VDD |  | A |  |  | 5 |  |  |  | 22AWG |
| DGND |  | b |  |  |  |  |  | 22 | 22AWG |
| FSYNC2 |  | D |  |  |  |  |  | 10 | 26AWG |
| LINE3 |  | T |  |  |  |  |  | 14 | 26AWG |
| LINE1 |  | C |  |  |  |  |  | 2 | 26AWG |
| READ2 |  | F |  |  |  |  |  | 12 | 26AWG |
| RESET3 |  | X |  |  |  |  |  | 13 | 26AWG |
| LSYNC4 |  | Z |  |  |  |  |  | 34 | 26AWG |
| LSYNC2 |  | L |  |  |  |  |  | 9 | 26AWG |
| PIXEL4 |  | a |  |  |  |  |  | 36 | 26AWG |
| FSYNC4 |  | c |  |  |  |  |  | 35 | 26AWG |
| FSYNC1 |  | E |  |  |  |  |  | 4 | 26AWG |
| LINE2 |  | B |  |  |  |  |  | 8 | 26AWG |
| READ3 |  | H |  |  |  |  |  | 18 | 26AWG |
| RESET4 |  | Y |  |  |  |  |  | 19 | 26AWG |
| RESET2 |  | J |  |  |  |  |  | 7 | 26AWG |
| LSYNC3 |  | N |  |  |  |  |  | 15 | 26AWG |
| PIXEL3 |  | S |  |  |  |  |  | 17 | 26AWG |
| FSYNC3 |  | V |  |  |  |  |  | 16 | 26AWG |
| LINE4 |  | U |  |  |  |  |  | 33 | 26AWG |
| READ4 |  | W |  |  |  |  |  | 37 | 26AWG |
| READ1 |  | G |  |  |  |  |  | 6 | 26AWG |
| RESET1 |  | K |  |  |  |  |  | 1 | 26AWG |
| LSYNC1 |  | M |  |  |  |  |  | 3 | 26AWG |
| PIXEL2 |  | P |  |  |  |  |  | 11 | 26AWG |
| PIXEL1 |  | R |  |  |  |  |  | 5 | 26AWG |

Notes

- Coax cables only connected at one end  as specified.

- Use thicker cable when specified.

- Connect CHASSIS to GND as specified.

- All 9,15 and 37 way D types are female plugs.

Assignments for  the bias voltages to the DAC channels are shown below.

VP1 BOARD DAC CHANNELS - P2

VOLTAGE	DAC		Pin

BIAS A		DAC1,1

BIAS B		DAC1,2

OFFSET4	DAC1,3		1

VRESET	DAC1,4		2

HIGH1		DAC2,1		3

BIASGATE	DAC2,2		4

VDD		DAC2,3		5

HIGH2		DAC2,4		6

VP2 BOARD DAC CHANNELS - P2

VOLTAGE	DAC		Pin

BIAS A		DAC1,1

BIASB		DAC1,2

HIGH3		DAC1,3		1

HIGH4		DAC1,4		2

OFFSET3	DAC2,1		3

OFFSET2	DAC2,2		4

OFFSET1	DAC2,3		5

BIAS PWR	DAC2,4		6

Cable from SDSU Controller to Preamp Box

| 50 way MALE PLUG | SIGNAL NAME | CABLE TYPE | 50 way FEMALE PLUG |
| --- | --- | --- | --- |
| 1 | +15V supply | 22AWG | 1 |
| 34 | -15V supply | 22AWG | 34 |
| 19 | ANALOG GND | 22AWG | 19 |
| 39 | HIGH1 | 26AWG | 39 |
| 23 | HIGH2 | 26AWG | 23 |
| 7 | HIGH3 | 26AWG | 7 |
| 24 | HIGH4 | 26AWG | 24 |
| 41 | VRESET | 26AWG | 41 |
| 18 | OFFSET1 | 26AWG | 18 |
| 36 | OFFSET2 | 26AWG | 36 |
| 37 | OFFSET3 | 26AWG | 37 |
| 38 | OFFSET4 | 26AWG | 38 |
| 50 | BIASGATE | 26AWG | 50 |
| 17 | BIASPOWER | 26AWG | 17 |
| 22 | BIAS SHIELD | COAX SHIELDING | ONE SHIELD OVER ALL THE ABOVE ONLY CONNECT MALE END |
| 2 | OUT1 | COAX | 2 |
| 21 | OUT1 SHIELD |  | ONLY CONNECT AT MALE END |
| 3 | OUT2 | COAX | 3 |
| 21 | OUT2 SHIELD |  | ONLY CONNECT AT MALE END |
| 4 | OUT3 | COAX. | 4 |
| 21 | OUT3 SHIELD |  | ONLY CONNECT AT MALE END |
| 5 | OUT4 | COAX | 5 |
| 21 | OUT4 SHIELD |  | ONLY CONNECT AT MALE END |
| 20 | OVERALL SHIELD | COAX SHIELDING | ONE SHIELD OVER ALL CABLES - ONLY CONNECT MALE END |

Clock cable connection between two AMP16-26 connectors

| Signal Name | Amphenol 16-26 free plug |  | Amphenol 16-26 free socket | Comments |
| --- | --- | --- | --- | --- |
| VDD | A |  | A | 22AWG |
| DGND | b |  | b | 22AWG |
| VDD SHIELD | b |  | n/c |  |
| PIXEL1 | R |  | R | 26AWG |
| LSYNC1 | M |  | M | 26AWG |
| FSYNC1 | E |  | E | 26AWG |
| READ1 | G |  | G | 26AWG |
| RESET1 | K |  | K | 26AWG |
| LINE1 | C |  | C | 26AWG |
| CLOCK SHIELD1 | b |  | n/c |  |
| PIXEL2 | P |  | P | 26AWG |
| LSYNC2 | L |  | L | 26AWG |
| FSYNC2 | D |  | D | 26AWG |
| READ2 | F |  | F | 26AWG |
| RESET2 | J |  | J | 26AWG |
| LINE2 | B |  | B | 26AWG |
| CLOCK SHIELD2 | b |  | n/c |  |
| PIXEL3 | S |  | S | 26AWG |
| LSYNC3 | N |  | N | 26AWG |
| FSYNC3 | V |  | V | 26AWG |
| READ3 | H |  | H | 26AWG |
| RESET3 | X |  | X | 26AWG |
| LINE3 | T |  | T | 26AWG |
| CLOCK SHIELD3 | b |  | n/c |  |
| PIXEL4 | a |  | a | 26AWG |
| LSYNC4 | Z |  | Z | 26AWG |
| FSYNC4 | c |  | c | 26AWG |
| READ4 | W |  | W | 26AWG |
| RESET4 | Y |  | Y | 26AWG |
| LINE4 | U |  | U | 26AWG |
| CLOCK SHIELD4 | b |  | n/c |  |

INGRID CRYOSTAT INTERNAL  CONNECTIONS

HAWAII

## INGRID CRYOSTAT

## RADIATION SHIELD

## DIGITAL

## ANALOGUE

# AMPHENOL 16-26 -> FANOUT BOARD J2 (MDM37) CONNECTOR

| Signal Name | MDM-37 (J2) | Amphenol 16-26 | Comments |
| --- | --- | --- | --- |
| VDD | 21 | A | copper |
| DGND | 19 | b | copper |
| DGND | 20 | b | copper |
| VDD SHIELD | 22 | n/c | shield these wires |
| PIXEL1 | 18 | R | manganin |
| LSYNC1 | 13 | M | manganin |
| FSYNC1 | 36 | E | manganin |
| READ1 | 31 | G | manganin |
| RESET1 | 26 | K | manganin |
| LINE1 | 8 | C | manganin |
| CLOCK SHIELD1 | 14 | n/c | shields these wires |
| PIXEL2 | 17 | P | manganin |
| LSYNC2 | 12 | L | manganin |
| FSYNC2 | 35 | D | manganin |
| READ2 | 30 | F | manganin |
| RESET2 | 25 | J | manganin |
| LINE2 | 7 | B | manganin |
| CLOCK SHIELD2 | 9 | n/c | shields these wires |
| PIXEL3 | 16 | S | manganin |
| LSYNC3 | 11 | N | manganin |
| FSYNC3 | 34 | V | manganin |
| READ3 | 29 | H | manganin |
| RESET3 | 24 | X | manganin |
| LINE3 | 6 | T | manganin |
| CLOCK SHIELD3 | 4 | n/c | shields these wires |
| PIXEL4 | 15 | a | manganin |
| LSYNC4 | 10 | Z | manganin |
| FSYNC4 | 33 | c | manganin |
| READ4 | 28 | W | manganin |
| RESET4 | 23 | Y | manganin |
| LINE4 | 5 | U | manganin |
| CLOCK SHIELD4 | 32 | n/c | shields these wires |

FANOUT BOARD J1 -> HERMETIC CONNECTOR

| Signal Name | MDM-31 (J1) | Amphenol 20-41 | Comments |
| --- | --- | --- | --- |
| OUT1+ | 13 | S | twist with OUT1- |
| OUT1- | 29 | R | twist with OUT1+ |
| OUT1 SHIELD | 12 | n/c | shield OUT1 |
| OUT2+ | 11 | P | twist with OUT2- |
| OUT2- | 27 | g | twist with OUT2+ |
| OUT2 SHIELD | 10 | n/c | shield OUT2 |
| OUT3+ | 9 | N | twist with OUT3- |
| OUT3- | 25 | f | twist with OUT3+ |
| OUT3 SHIELD | 8 | n/c | shield OUT3 |
| OUT4+ | 7 | M | twist with OUT4- |
| OUT4- | 23 | e | twist with OUT4+ |
| OUT4 SHIELD | 6 | n/c | shield OUT4 |
| HIGH1 | 16 | r | manganin |
| HIGH2 | 15 | h | manganin |
| HIGH3 | 18 | T | manganin |
| HIGH4 | 17 | L | manganin |
| BIASGATE | 4 | U | manganin |
| BIASPOWER | 3 | j | manganin |
| VRESET | 2 | q | manganin |
| AGND | 30 | d | copper |
| AGND | 29 | d | connected together |
| AGND | 26 | d | connected together |
| AGND | 24 | d | connected together |
| AGND | 22 | n/c |  |
| AGND | 14 | n/c |  |
| AGND | 5 | n/c |  |
| BIAS SHIELD | 31 | n/c | shield these wires |
| TGND | 20 | n/c | not used |

Notes

- Each of the four outputs from the HAWAII device is a shielded twisted pair. Copper wire (PTFE coated) should be used for the lines and the miniature co-ax cable shield wire should be stripped and used to screen the twisted pairs. The screen should only be connected at one end as specified.

- Copper wire should be used for the AGND connections. These wires are paralleled up as specified.

- An outer shield should be connected over the Bias lines as specified. This should only be connected at one end.

- Use manganin where specified.

- Each sub set of clock lines is to be shielded. The technician will need to strip co-ax shield cable from the miniature co-ax cable supplied and use this to shield the clock groups as specified. The shield must only be connected to the fanout board end. It must NOT be connected to the hermetic connector end.

- All cables should be made as short as practically possible.

- These cables may have to go via a connector block on the radiation shield. If so then all connections should be straight through including the shields which must go via pins in the connectors and not via the connector shells. This means that the connectors on the radiation shield may be of a different size and format than the connectors at either end of the cable.

- Connections are also required to be made to the J3 MDM9 way connecotr on the fanout board. This connector allows control of a preflash LED on the board an temperature checking of the board. The board connections are  as shown below. These wires must go to another hermetic 16-26 Amphenol on the cryostat body. This connector’s connections will be shared with other temperature sensors’ wires ans micro-switches.

DIODE TEMPERATURE SENSE/PREFLASH - J3 –SOCKET PCB MOUNTED 90 DEGREES  MALE SOCKET MDM-9PCBR

| PIN | SIGNAL |
| --- | --- |
| 1 | TEMP DIODE+ |
| 2 | TEMP DIODE- |
| 3 | PREFLASH LED+ |
| 4 | PEFLASH LED- |

INGRID SAFETY EARTH CONNECTIONS

# PSU

# CRYOSTAT

# TELESCOPE

# SDSU

# CCC HEAD

# COMPRESSOR

# PREAMP

### INGRID GROUNDING PLANS

- SEPARATE AGND and DGND supplied from the SDSU Controller.

These come to the “star” point in the controller but are supplied separately to the HAWAII device fanout board. This is done to reduce clock feedthrough onto the analog side electronics.

- The clocks and biases/outputs are supplied on separate cables to the HAWAII device. Each set of signals is shielded with one end of the shield connected to a ground connection at the fanout board end but disconnected at the other end. The shield connection is taken through separate pins on the connectors in the radiation shield.

- The cryostat body should be isolated from the telescope structure. The preamp box and SDSU controller are mounted off the cryostat using electrically insulated supports. The cryostat does not get its safety earth connection through this path but via the Closed Cycle Cooler Head. If  the CCC is not used and disconnected then a safety earth connection must be made direct to the cryostat body.

- The cables for the motors should be wired up as per Berger Lahr instructions. This means that the cables ought to be shielded with the shield only being connected at the drive end. internal cryostat shielding will not be required.