From pcm@ing.iac.esTue Dec 17 15:05:03 1996

Date: Tue, 17 Dec 1996 13:30:40 +0000 (GMT)

From: Peter Moore <pcm@ing.iac.es>

To: kmd@ing.iac.es

Cc: awr@ing.iac.es, cwmj@ing.iac.es

Subject: Whircam testing

Below are the results of a series of tests performed on whircam to try

and derive a cause for the charring of constantine wires in the clock

harness in the detector housing.

These tests were based on the assumed ohmic heating of all or some of

the 18 conductors sufficient to cause complete loss of the enamel

insulation and charring (through evaporation) of the silk protection thread.

The tests were carried out in an attempt to discover the source of the

current by creating fault scenarios and then trying to eliminate these

possible causes.

1.0 Starting with the known, final conditions we have:

1.1 Whircam failed when it was noticed that the images in movie mode

where no longer produced. No whircam error was reported by the system;

cooling, electronics, software were all shown as normal.

1.2. At the time of the failure, the telescope was low on the horizon and

was in the process of instrument rotation limits testing. After

the failure the telescope was moved to zenith.

1.3 The instrument was left in normal mode for a period of at least 4

hours after the fault occurred, at which time the detector temperature

had risen from nominally 30K to 114K and the cold shield from 70K to 85K.

At this point the instrument was shutdown normally.

1.4 After dismounting whircam and during subsequent inspection it was found

that the constantine portion of the 18 conductors in the array clocks cable

leading from the optical table thrust plate to the detector array pcb

where charred black. These wires did not appear to be shorted to either the

detector housing or to each other.

1.5 Two bond wires on the InSb array where also found to be lifted from

their bond pads (Fast phase 1 clock and fast sync). They did not show the

normal ball ends that occur from fusing through over current.

1.6 It was observed that the clearance hole in the optical table was small

enough to allow the second stage cold finger too touch the optical table.

Suppositions that can reasonably be derived from conditions 1.3 and 1.4. are:

1.7 That the fault is persistent

1.8 That the damage caused by heating in the array clocks cable was the

cause for the elevated temperatures observed before shutdown.

2.0. Electronic Testing

2.1. Conclusions

Without considering influences external to whircam, the only sources of

current that could be dissipated through the array clock cable are from the

clock drives themselves, the bias supplies, or the preamp sub assy. All

of these have been tested and found functional.

Considering influences external to whircam, the possibility of causing a

potential difference between the cryostat and electrical ground exists,

however, the path of any current delivered would not include the array

clock cable. It would flow in the low impedance paths of the cable shields.

No cause of the damage observed to the array clock cable can be found at

this stage.

We await the report from SBRC detailing the condition of the array before

attempting any further analysis.

2.2 Testing

2.2.1 Fault Scenario 1

That the cold finger did touch the optical table causing a current to

flow, from a difference in potential of the cryostat with respect to

the alice clock ground, through the cold finger to the array to the array

pcb and to the alice controller clock board ground.

2.2.2 Modifications to the ground wiring that tie the alice / whircam cable

shields together at the preamp box assure that there can be no difference

of potential between cryostat and alice.

2.2.3 Although it is likely that the cold finger touched the optical table

due to the displacer sagging at the telescope altitude when the fault

occurred, and this would lead to the array warming up, this condition

would not have persisted when that telescope was driven to zenith. This

would have removed the heat load and so cannot explain the elevated

temperatures found at shutdown.

2.2.4 Fault Scenario 2

That one or several clock drivers in the alice controller failed causing,

at worst case, +15v DC drive to the array. This potential would lead to

a current flow from the alice clock driver, through a 91 ohm resistor, through

the (approx) 1 ohm constantine portion of the clock cable, on to the array

pcb, through a zener diode, back through the constantine ground return, to the

power supply ground of the clock driver board.

2.2.5 The max current to flow in this scenario is 170 ma which would have

caused over 2.5 watts to be dissipated in the 91 ohm resistor. This would

have caused destruction of this resistor which is not the case.

2.2.6 The clock drives where checked and verified to be their correct values.

2.2.7 Fault Scenario 3

That one or more of the bias supplies to the array became, at worst case,

+15v thus forward biasing the protection zener on the array pcb and feeding

current to the clock ground via the clock cable ground portion of the array

clocks cable. (this is possible via the modification to the grounding scheme)

2.2.8 The bias supplies pass through a low pass filter on the array pcb which

requires that the 170 ma current would pass through a 100 ohm resistor, thus

distroying this. This was not apparent.

2.2.9 The ground return conducters in the array clocks cable are in parallel

so the load would be shared by 9 separate conducters thus minimizing any

ohmic heating.

2.2.10 The bias voltages were checked with the alice controller functioning

and found to be at their correct values.

2.2.11 Fault Scenario 4

That the normal dynamic load on the clock drivers creates ohmic heating in

the constantine portion of the array clocks cable which cannot be thermally

dissipated via conduction to the array pcb. This would lead to a

temperature rise of sufficient degree over time to cause degradation of the

wire insulation enamel and silk.

2.2.12 A numerical analysis was done to determine if a short circuit current

of 170ma (worst case clock driver failure) could damage the wire. It was

found that for a simple model of pure thermal conduction through a 5cm length

of constantine wire a temperature rise of 614K can be achieved in 4

hours, sufficient to cause the observed damage. This is a dissipation of

approx 30mw which puts this finding is in some conflict with the

dissipation measured in the next paragraph.

2.2.13 A destructive test was carried out on a representative portion of

constantine wire to determine the temperature and dissipation required to

cause the degree of damage observed. The temperature reached was approx

570K to achieve charring in free air and approx 300 mw in vacuum for the

beginning of degradation to occur. The normal dissipation was calculated

for a 0.5 sec movie mode readout and found to be approx 7nw considering 100pf

pmos gate capacitance which is magnitudes short of the required value

(this value should be checked !).

To the best of our knowledge all auxillary circuits (temp monitoring,

displacer drive electronics, filter wheel drives, etc) have been

tested for galvanic isolation from the array electronics.

===+===+===+===+===+===+===+===+===+ ING +===+===+===+===+===+===+===+===

=   Peter Moore, Isaac Newton Group, La Palma Observatory, Spain.       =

=   E-mail pcm@ing.iac.es                  Voice Office +34 22 405566   =

=   S-mail Apartado de Correo 321,           Fax Office +34 22 405646   =

=          38780 Santa Cruz de La Palma,     Voice Home +34 22 423945   =

=          Canary Islands, Spain.                                       =

===+===+===+===+===+===+===+===+===+ LPO +===+===+===+===+===+===+===+===