## Noise  Free  CCDs’

Peter Moore – Isaac Newton Group – August 1999.

# summary

An analysis of the degradation caused by electronic noise contamination in astronomic observations using charge-coupled devices is presented. From this a technique for removing the electronic noise component from such observations is developed such that the SNR of observational measurement is improved.  In addition the technique provides for increased dynamic range, reduced sensitivity of the electronics to temperature effects, and de-couples electronic conversion gain from readout speed.

# 1. introduction

Given the extensive use of charge-coupled devices (CCDs) to detect photons in astronomy, the limitations of these devices still prevail to limit the science that can be obtained from them. The criteria of any planned observation must take into account the instrumental limitations to enable sufficient  signal to be derived for the measurement accuracy required. In specific areas of investigation, for example time resolution photometry and faint object spectroscopy, observations are compromised by a lack of strong signal that, when coupled to the limitations of the overall instrument throughput, make observation or interpretation impossible.  In instrumentation for astronomy, two characteristics of a CCD detector and readout electronics contribute directly to the available SNR of any observation. These are the quantum efficiency of the detector at the desired wavelength Q and the electronic or ‘readout’ noise Nrd associated with processing of the photon generated signal from the detector. Considering these two characteristics of the CCD a model of expected SNR S can be derived given the incoming photon flux from the object of interest  I, the photon flux from background sky B, the time of integration t, and the number of pixels illuminated by the object of interest n. Thus the SNR is given by

S = Q It / [Q(I + B)t + nNrd2]1/2			(1)

The contribution that readout noise plays in SNR can be seen by plotting the relative SNR for three observational scenarios. These are plotted in the attached figures with SNR shown as a function of seeing (fig 1), as a function of integration time (fig 2), and as a function of object photon flux (fig 3).

From these illustrations it is seen that readout noise plays a significant role in the determination of observation SNR under all conditions. Under marginal conditions, noise free observing has particular advantage in improving observation quality. In such cases noise free observing could make the difference between success and failure of the science objectives. However, it is to be noted that all observations will benefit productively from noise free observing because, as readout noise is independent of the observed photon flux , readout noise sets the lower threshold of detection for the instrument.

# 2. defining noise

Fig 4 shows a typical CCD based signal processing chain and defines the types of noise present and their origin within the signal processing chain. These noise types can be broadly classified into 3 classes described by their origin. Intrinsic noise sources are those generated by the process of detecting and conditioning the expected signal. Examples of this class are dark current, 1/f noise, quantization error, etc. The second class include sources of noise that are man made and influence the detection process by their proximity without actually taking part in the process itself, for example EMI.  The third class is those sources generated by natural causes, for example, cosmic rays, sky brightness, lightning, etc. It is customary to specify all noise sources in units of equivalent signal (electrons) referenced to the detector.

PRNU, dark current, and fixed pattern noise are intrinsic noise sources that are proportional to signal or time. Methods exist so that these contributions can be reduced to sub-electron levels or quantified and modelled with sufficient accuracy to remove their effect from science images. This paper will concentrate on the electronic subset of the intrinsic class, and develop methods to eliminate the components of reset, Johnson, 1/f, quantization, and drift noise sources. These components are independent of signal and they influence all signals in a similar manner, including signal and time dependent noise sources such PRNU and dark current. Because of this, the electronic noise contribution effectively sets the lower limit of sensitivity by masking low level signals under the electronic noise component. These noise sources, with the possible exception of drift, produce noise that manifests itself as pixel to pixel variation. Individual noise sources add in quadrature to form a total noise power for a given system. The ideal CCD camera system would remove these noise components completely or reduce their combined amplitude to below one equivalent photon (i.e. sub-electron noise).

Consider the character of these electronic noise components. Reset noise is the ambiguity in the charge deposited into the reset well before signal is clocked in. It is caused by the reset ripple voltage and Johnson noise current of the reset FET resistance. This latter component is a thermal noise source in parallel with the reset node capacitance. The reset contribution, Nrst, can be expressed as

Nrst  = [(kTCB )1/2 + CVripple] / q    (e- rms)		(2)

Where k = Boltzmann constant, T = temperature, C = the capacitance value of the reset node, B = bandwidth of system, Vripple =  rms reset ripple voltage, and q = electron charge.

Johnson noise has a characteristic ‘white’ frequency spectrum i.e. the noise power density for any unit bandwidth is constant from DC to infinity. This implies that the white noise component of any electronic system is proportional to the bandwidth of the system. Hence the inclusion of system bandwidth in the above equation. 1/f noise sources however, have a characteristic power that is proportional to current density and approximately inversely proportional to frequency. This implies that high frequency systems contain a lower noise component from this source. The detector output amplifier and all gain components within the signal conditioning chain contribute both Johnson and 1/f components. The Johnson noise component of a signal chain element is given by

Njohnson = (4RCkT/Gq)B1/2 	(e- rms / Hz)		(3)

Where R is the signal path resistance, C = reset node capacitance, k = Boltzmann constant, T = temperature, G = signal gain, q = electron charge, and B is system bandwidth. A representative curve is shown in figure 5.

The origin of 1/f noise is not fully understood, however the component can be approximated by the empirical expression

N1/f = GIk B/Faq	(e- rms)			(4)

Where G = gain of the element, I = current density in the conduction path, k = empirical constant close to 3, B = system bandwidth, F = pixel read rate, a = empirical constant close to 2 for semiconductors, q = electron charge. Figure 6 show a representative curve for 1/f noise power.

Quantization noise will be generated by camera systems where the conversion process of the analog signal to a digital value has a resolution of less than one half of one equivalent photon (e-). This is most often the case where large dynamic range is required for sky fields containing bright objects. Since the conversion process has a limited dynamic range, usually 216 codes, the least significant bit ends up representing multiple photons and the signal becomes undersampled. When this occurs the signal amplitude acquires a noise component equal to

Nadc = (C / Gq) (VLSB / 121/2)	(e- rms)		(5)

Where C = reset node capacitance, G = overall signal gain, q = electron charge, and VLSB = voltage equivalent of the least significant converter bit.

In addition to quantization error, gain and offset error associated with non-perfect signal path elements and data converters are additional, fixed signal error sources but do not contribute a time or signal dependent change except by drift as described next.

Drift is characterized by being proportional to the operating temperature of the system and has a time constant much greater than the other noise components. However, given large detector array sizes, the amplitude of this component can become significant in long read times. Indeed, signal degradation due to drift is a serious problem when coherent data is required, spanning for example, one or more complete nights of observation. Currently, this contribution can only be removed by careful and frequent observations of standard metrics. Drift therefore may be modelled as

Ndrift = (CVdrftt)/(Gq)	(e- rms)		(6)

Where C = reset node capacitance, Vdrft = voltage drift per degree k, t = delta temperature change, G = overall signal gain, and q = electron charge.

Thus, a reasonable model for the inherent independent electronic noise component of any observation is given by

Nrd  = (Nrst2 + Njohnson 2 + N1/f 2 + Nadc 2 + Ndrift2 )1/2	(e- rms)	(7)

# 3. noise removal

Currently operating CCD cameras are designed with a compromise between medium frequency (maximise pixel rate) systems demanded for large CCD arrays that require high bandwidth (i.e. white noise limited performance) and ‘low noise’, low frequency low pass signal averaging systems dominated by 1/f noise. The classical method of CCD signal conditioning to eliminate noise components is to use a technique called correlated double sampling (CDS). This technique employs an integrator or low pass filter to integrate equal samples of the reset and signal values and derive an amplified difference signal. This technique is very effective in removing reset and 1/f components, however, the high frequency roll-off of the filter formed by the CDS integrator is determined by the pixel rate required.  The high pixel rates that are now demanded means that the choice of electronic components fast enough to accommodate the signal accurately becomes critical. This generally results in pushing the system bandwidth out wider than desired and allowing white noise components to become dominant. Figure 7 illustrates this.

The proposed technique addresses these problems by replacing the classic analog signal conditioning electronics with a software CDS thus reducing the complexity of analog circuitry in the front end. The front end will have a wide signal bandwidth and allow a very high analog to digital sample rate to be sustained on the signal. The analog signal conditioning is limited to a fixed gain stage, level shift, and anti-alias filtering to the Nyquist frequency determined by the sample rate. The sample rate will be in the order of 40 mega-samples / sec. The processing of this over-sampled signal is then done with firmware in a digital signal processor. The advantages of this technique are

- Once the signal is sampled, no noise is added by the signal processing.

- Much simpler analog processing allowing easier drift compensation.

- Conversion gain of system (e- / ADU) independent of readout speed.

- Digital signal processing allows much tighter control of processing design.

The combined application of these advantages is expected to allow the intrinsic electronic noise component to be reduced to below the equivalent one photon level, leaving the signal and signal dependent noise components available to the user. This is achieved by reducing elements contributing noise, stabilising the signal chain, and limiting the bandwidth to the minimum required to extract and process the signal. Figure 8 shows the sharpness of filter cut-off available with digital signal processing.

It is only recently that the technology has become available to realise this technique.  Even so, the fast, high-resolution analog to digital converters that are required by this technique will need to be built from lower resolution sub components now available commercially.

# 4. physical implementation

The technique described can be implemented so as to relate to existing ING and SDSU CCD controller architectures. This will result in a board level change to the existing equipment to bring this novel concept into operational use. This has the following advantages

- The technique can be introduced in a step by step manner with minimum disruption to existing equipment.

- The technique can be retrofitted to only those cameras that will benefit most from noise free observing, thus reducing cost.

- There will always be a fallback option if problems occur to the new equipment.

- The technique can be offered as a retrofit option to other users of the same equipment.

# 5.  expected performance

| Characteristic | Value | Units |
| --- | --- | --- |
| Signal channels per board | 2 |  |
| Max readout speed per channel | 1 | Mpix / sec. |
| Conversion gain | 1 | e- / ADU. |
| Conversion dynamic range | 218 | e- |
| Fixed gain steps | 1,2,4 | e- / ADU |
| Board noise contribution | 0 | e- |
| Combined detector / board electronic noise | < 0.5 | e- rms |
| Differential non linearity | < 0.25 | e- |
| Integral non linearity | < 1 | e- |
| Gain drift | 0 | e- |
| Offset drift | < 2 | e- 0 < t < 35 |

# 6. design implimentation

Fig 9 shows the basic architecture of a design to implement the described technique. The output of the CCD amplifier is resistively loaded to –5 volts to allow sufficient headroom for the difference amp and provide an independent quiet supply for the substrate reference. An offset voltage drive is provided by a D / A converter to subtract the mean reset offset voltage of the CCD amplifier. This is achieved in the low gain difference amplifier buffering the output signal and providing common mode rejection and impedance matching to the dc coupled second stage gain section. This section sets the overall sensitivity at ½ e- for the system. The signal is now presented to the low and high order stages where conversion to a digital value takes place. The low order stage samples the signal with full dynamic range upon the convert command from the sequencer, usually after reset settling time. The value thus obtained is fed back to the

7. costs and time scales

Table 1 Estimated costs to produce one system

| Task or Process | Labour time | Contract Costs | Part costs |  |
| --- | --- | --- | --- | --- |
| HW design | 3 wk |  |  |  |
| SW design | 4 wk |  |  |  |
| SW simulation | 2 wk |  |  |  |
| HW prototype | 2 wk |  | 2k |  |
| HW layout |  | 2k |  |  |
| Manufacture |  | 1k | 0.5k |  |
| Testing & characterisation | 2 wk |  |  |  |
| Documentation | 1 wk |  |  |  |
| TOTALS | 14 wks | 3k | 2.5k | 14 wks / 5.5k |

Table 2 Estimated costs to retrofit 10 systems

| Task or Process | Labour time | Contract costs | Part costs |  |
| --- | --- | --- | --- | --- |
| Manufacture |  | 2k | 5k |  |
| Test & Install | 3 wk |  |  |  |
| TOTALS | 3 wk | 2k | 5k | 6 wks / 7k |

8. figures

Fig 1. Effect of seeing on SNR.

Fig 2. Effect of Integration time on SNR.

Fig 3. Effect of Object photon flux on SNR.

Fig 4. Noise source model for CCD Camera system

Fig 5. Johnson noise component as a function of system bandwidth.

Fig 6. 1/f noise component as a function of pixel read rate.

| TBD | TBD |
| --- | --- |
| Fig 7. Bandwidth of current CCD controller signal path. | Fig 8. Bandwidth of proposed signal processing path. |

Fig 9. Architecture of proposed signal processing path.

9. references

(1)	J.R. Allington-Smith & H.E. Schwarz (1984)

Image photon detectors for optical astronomy

- James Janesick, Tom Elliot

History and advancements of large area array scientific CCD imagers.

- Gerald C. Holst

CCD arrays cameras and displays.

- Prof. D.A. Bell

Noise in electronic systems.

- James W. Beletic & Paola Amico

Optical detectors for astronomy.