Isaac

Newton

Group

Red

Imaging

Detector

SDSU INTERFACE CONTROL DOCUMENT

VERSION 4.0

Peter Moore 25th May 2002.

##### Version Summary

This version enhances the functionality of version 3.2 and removes the interface requirements for ImpB data acquisition i.e this version is only UltraDas compliant. Improvements are: Transfer of all timing features to the utility board via a call back scheme; Addition of the GRB command to allow fast frame grabs; Implementation of fast and slow readout schemes based on a generic change to the video pipeline settling time.

## Introduction

## This document describes the protocol used for communication between the Host computer system and the SDSU detector controller that forms part of the INGRID instrument. These sub systems are linked together using two fibre optic cables connected directly between the interface card in the Host Computer and the Timing Board of the SDSU controller. The uplink from Host to SDSU is a slow link operating at 4 MHz  which is used for sending commands to the SDSU controller. The downlink operates at 50 MHz and is used for sending responses to the uplink commands and also image data to the Host System. There is also an electrical communication serial link operating between the Timing Board and Utility Board in the SDSU controller using the standard Motorola DSP SSI interface.

This document presents the protocol in a table format. The significance of each of the columns is as follows:-

Column 1 = Executable Command

Column 2 = Originator of Command

Column 3 = Destination of Command (specifies if available after BOOT or only when an application APPL has been uploaded)

Column 4 = Number of words in command

Column 5 = Response to the command

Column 6 = Specifies the action taken on receiving the command

Column 7 = Remarks and more information

The format of the messages sent between HOST and SDSU has been described many times elsewhere; suffice to say that each command or response consists of 2 – 7 words. Each word is made up of 33 bits. Of these, 24 bits are valid and the rest are used as header information. The programmer need not worry about this header information. The hardware strips away the header information to leave the expected 24 bit word, which is then processed by the SDSU controller.

The downlink is used for responses and image data. When transmitting image data, then it takes the form of 17 bit words, with one stop bit and 16 data bits. The host programmer should know when to expect the 33 bit packet or 17 bit packet and act accordingly. Again this is described in more detail in the SDSU documentation.

The intended audience for this document are those who are programming either the Host computer end or SDSU controller end of the fibre links.

This document also gives a table of MailBox addresses in the SDSU controller. These locations are used to communicate variables between host and sdsu and sdsu and host. For this purpose the X: memory area is deemed writable by the host and the Y: memory space readable from the host: The reverse is true for the sdsu controller.

| ### Command | #### Source | Destin | Words | # Response | Action | Remarks |
| --- | --- | --- | --- | --- | --- | --- |
| TDL nnnnnn 0 nnnnnn  ffffff (BOOT) | HOST | TIMING,UTILITY | 3 | nnnnnn | Test Data Link. Destination echoes nnnnnn back to Source. |  |
| RDM maaaaa 0  aaaaa  0ffff (BOOT) | HOST | TIMING, UTILITY | 3 | dddddd | ReaD Memory. Read DSP address maaaaa. Returned data = dddddd. The most significant nibble of the address indicates the memory type. m = 1: P memory m = 2: X memory m = 4: Y memory m = 8: EEPROM | This command is used to read memory locations for low level fault finding or checking the simple variables, e.g. elapsed integration time. |
| WRM maaaaa dddddd 0  aaaaa  0ffff 0 dddddd  ffffff (BOOT) | HOST | TIMING, UTILITY | 4 | DON | WRite Memory. Write to DSP address maaaaa with data = dddddd. The most significant nibble of the address indicates the memory type. m = 1: P memory m = 2: X memory m = 4: Y memory m = 8: EEPROM | This command is used to write to memory locations for low level fault finding and downloading the application code. |
| LDA n n = 0 … 256 (BOOT) | HOST | TIMING UTILITY | 3 | ERR | Support for application program load from EEProm. Not implemented and will always return ERR | Redundant / future command. Applications are downloaded from host only. |

| ### Command | #### Source | Destin | Words | # Response | Action | Remarks |
| --- | --- | --- | --- | --- | --- | --- |
| RDC (APPL) | HOST | TIMING | 2 | IMAGE DATA, DON | Performs one readout of array and sends data to host. |  |
| GRB (APPL) | HOST | TIMING | 2 | IMAGE DATA, DON | As per MRA but does not do the array flushing sequence so it is faster. |  |
| MRA n 0n 0xffffff (APPL) | HOST | TIMING | 3 | IMAGE DATA, DON | Execute Multiple Non Destructive Read (MNDR) consisting of an array reset, n Reads – Integration and n reads. transmits DON at completion | The type of data sent depends on the flag set by the DAT command. A DON command is sent after the image data is sent. |
| RUR n 0n 0xffffff (APPL) | HOST | TIMING | 3 | IMAGE DATA, DON | Execute Read Up Ramp readout consisting of an array reset, MNDR, sequence of n Integrations and MNDR’s. transmits DON at completion | This command is a place holder for future read up ramp readout scheme. Just returns DON. |
| ABR (APPL) | HOST | TIMING | 2 | DON | Abort Mode that Controller is in and return to Idle Mode – Resets READ/TIMER/VIDEO/IDLE mode flags. | ABR can be sent anytime after MRA command but will only be processed after post reset reads are completed. See v2.1 release notes for further info. |

| ### Command | #### Source | Destin | Words | # Response | Action | Remarks |
| --- | --- | --- | --- | --- | --- | --- |
| CLR (APPL) | HOST | TIMING | 2 | DON | Performs a complete reset (Clear) of the detector array. | Takes 1.3 ms + comms time to complete. |
| DAT n 0  n  3 (APPL) | HOST | TIMING | 3 | DON | Determines type of data that MRA command transmits where n equals 0 => array data, 1 => 1111,2222,etc. 2=>0,1,2, 3..65535. 3 => array data ‘read up ramp’ mode (not implimented as yet) | Image data is transmitted faster than the test data because its algorithms runs from fast DSP memory whereas the test data runs from slow memory Mode set to 0 by reset & CON |
| IDL (APPL) | HOST | TIMING | 2 | DON | Enables periodic idle mode clearing of detector array | This mode is enabled as default when the PON/CON command is issued. |
| STP (APPL) | HOST | TIMING | 2 | DON | Disables periodic idle mode clearing of detector array | Used to establish ‘INTEGRATE’ mode |
| LSP (APPL) | HOST | TIMING | 2 | DON | Sets a series of parameters for performing lower speed readouts | Set low speed readout mode. |
| HSP (APPL) | HOST | TIMING | 2 | DON | Sets a series of parameters for performing higher speed readouts | Set High speed readout mode |

| ### Command | #### Source | Destin | Words | # Response | Action | Remarks |
| --- | --- | --- | --- | --- | --- | --- |
| SBS (APPL) | HOST | TIMING | 2 | DON | Re-establishes bias voltages from table to hardware. | Allows discrete bias voltage values to be changed without rebooting the controller. Be Careful ! |
| CON (APPL) | HOST, UTILITY | TIMING | 2 | DON | Switch voltages ON to array Clears readout mode to 0 (real array data). | Automatically sent by the PON command |
| COF (APPL) | HOST, UTILITY | TIMING | 2 | DON | Switch voltages OFF to array | Automatically sent by the POF command. |
| PON (APPL) | HOST | UTILITY | 2 | DON | Enables +15V and –15V to analogue circuitry in controller + loads clock and video boards with voltage values. | This command now calls CON command in the Timing board i.e it powers up the complete system. Must be used before telemetry is read. |
| POF (APPL) | HOST | UTILITY | 2 | DON | Disable voltages to analogue circuitry and sets array to safe mode i.e biases and clocks clamped to gnd. | Calls COF in the timing board before powering down main supplies. |
| OSH (APPL) | HOST | UTILITY | 2 | DON | Open shutter | Remains OPEN until RESET or CSH sent |
| CSH (APPL) | HOST | UTILITY | 2 | DON | Close shutter |  |

| ### Command | #### Source | Destin | Words | # Response | Action | Remarks |
| --- | --- | --- | --- | --- | --- | --- |
| PFL (APPL) | HOST | UTILITY | 2 | DON | Turns on preflash LED, starts preflash elapsed time timer and eturns immediately. Preflash LED will be switched off after the elapsed time >= demanded preflash time | Contary to spec DAS-18. Note that this command returns immediately. This to allow preflash periods greater than the host command timeout period. |
| LON (APPL) | HOST | UTILITY | 2 | DON | Switch internal LED ON | Remains ON until RESET or LOF sent |
| LOF (APPL) | HOST | UTILITY | 2 | DON | Switch internal LED OFF. |  |
| BEX (APPL) | HOST, TIMING | UTILITY | 2 | DON | Begins exposure sequence by opening shutter and starting elapsed timer running. When elapsed time is >= to demanded time it will close shutter. Called by MRA, GRB, RUR | Sends DON flag after shutter opened but before integration time reached. Interacts with Timing board to effect complete exposure. Not normally needed to be sent from host. |
| PEX (APPL) | HOST | UTILITY | 2 | DON | Pause exposure. Closes shutter and stops elapsed timer running | Returns error if exposure not active |
| REX (APPL) | HOST | UTILITY | 2 | DON | Resumes exposure. Opens shutter and continues elapsed timer. | Returns error if exposure not active and/or paused by PEX command |

| ### Command/ Msg | #### Source | Destin | Words | # Response | Action | Remarks |
| --- | --- | --- | --- | --- | --- | --- |
| DEX (APPL) | HOST | UTILITY | 2 | (Delayed) DON | Detect end of exposure waits for maximum of 8 seconds for elapsed timer to be = to demanded time then returns status | Returns ERR if 8 seconds elapse before end of exposure. |
| CBK (APPL) | HOST, TIMING, UTILITY | TIMING, UTILITY | 2 |  | Internal message used to synchronize the timing and utility commands to perform an integration and readout | Can be simulated from the host but with zero purpose for operations. Remains available to host as a debug tool. |
| ERR (BOOT,APPL) | TIMING, UTILITY | HOST | 2 |  | Error message sent as result of command lookup error or incompatible command. |  |
| SYR (BOOT) | TIMING | HOST | 2 |  | Informs HOST system that SDSU controller has performed a PWR RESET. | required for IRCAM compatibility |
| DON (BOOT,APPL) | TIMING, UTILITY | HOST | 2 |  | Acknowledge message sent as result of successful command execution |  |

Notes to COMMAND Table :-

- Not all commands are available at all times. Column #1 indicates whether each command is: (i) a BOOT command which is available aftyer power-up or reset or when the application is loaded, or (ii) an APPLication command which is available only for a unique application program which has been downloaded from the Host system.

Appendix A

This table shows a typical flow of commands and responses after the system has been reset.

| Sequence | HOST command | SDSU Response | Description |
| --- | --- | --- | --- |
| 1 |  |  | System Reset |
| 2 |  | SYR | SDSU replies that it has RESET |
| 3 | 000203 TDL 555555 | 020002 555555 | Test the link to the TIMING board |
| 4 | 000303 TDL AAAAAA | 030002 AAAAAA | Test the link to the UTILITY board |
| 5 | 000203 RDM 100006 | 020002 xxxxxx | Read version no. of Timing board boot code |
| 6 | 000303 RDM 100006 | 030002 xxxxxx | Read version no. of Utility board boot code |
| 7 | *.lod file downloaded using WRM command |  | Download Timing Board Application Code |
| 8 | *.lod file downloaded using WRM command |  | Download Utility Board application code |
| 9 | 000302 PON | 030002 DON | Switch supplies ON to boards |
| 10 | 000203 RDM 1001FE | 020002 001FF0 | Read NBAX Timing board Mailbox location |
| 11 | 000203 RDM 1001FF | 020002 003FF0 | Read NBAY Timing board Mailbox location |
| 12 | 000303 RDM 1001FE | 030002 0000F0 | Read NBAX Utility board Mailbox location |
| 13 | 000303 RDM 1001FF | 030002 0000F0 | Read NBAY Utility board Mailbox location |
| 14 | 000304 WRM 0200F0 0007D0 | 030002 DON | Set integration time to Utility Mailbox area. |
| 15 | 000203 MRA 1 | Image Data…. 020002 DON | Sends image data then DON back |

Appendix B

Addresses which can be accessed using the RDM command.

| Board | Address | Available | Format | Description |
| --- | --- | --- | --- | --- |
| Timing / Utility | P:7 | BOOT | Binary | Version No. of Application code |
| Timing / Utility | P:1FE | BOOT | Binary | Noticeboard offset into X:mem space for UltraDas (NBAX) |
| Timing / Utility | P:1FF | BOOT | Binary | Noticeboard offset into Y:mem space for UltraDas (NBAY) |
| Timing | X:0 | BOOT | Boolean | Timing status word (note 1). |
| Timing | X:22 | APPL | Integer | Readout mode |
| Timing | X:24 | APPL | Integer | Number of columns in image data |
| Timing | X:25 | APPL | Integer | Number of rows in image data |
| Timing | X:26 | APPL | Integer | Number of reads in last MRA sequence |
| Timing | X:27 | APPL | Integer | Number of pre-readout detector reset cycles |
| Timing | X:28 | APPL | Integer | Number of read precondition (flush) cycles |
| Timing | NBAX | APPL | Int table | Window skip/read table origin |
| Timing | NBAX+FC | APPL | Boolean | UltraDas Active Flag |
| Timing | NBAX+FD | APPL | Integer | Binning Factor in x (Not used !) |
| Timing | NBAX+FE | APPL | Integer | Binning Factor in y (Not used !) |
| Timing | NBAX+FF | APPL | Boolean | Window table enable flag |
| Timing | NBAY | APPL | Integer | Error number (not used !) |
| Timing | NBAY+1 | APPL | Integer | Window table dimension |
| Timing | NBAY+2 | APPL | Boolean | Readout speed |
| Timing | NBAY+3 | APPL | Integer | Pixel rate |
| Timing | NBAY+4 | APPL | Boolean | Clock sequencer status |
| Timing | NBAY+5 | APPL | Integer | Number of readouts since boot |

| Board | Address | Available | Format | Description |
| --- | --- | --- | --- | --- |
| Utility | X:0 | BOOT | Boolean | Utility status word (note 2). |
| Utility | Y:0 | BOOT | Byte Int | Detector ID code (1Eh) |
| Utility | NBAX | BOOT | Integer | Demanded exposure time (ms) |
| Utility | NBAX+1 | BOOT | Integer | Demanded detector temp (mk) |
| Utility | NBAX+1 | BOOT | Integer | Demanded preflash time (ms) |
| Utility | NBAY | BOOT | Integer | Elapsed exposure time (ms) |
| Utility | NBAY+1 | BOOT | Integer | Current detector temp (mk) |
| Utility | NBAY+2 | BOOT | Integer | Elapsed preflash time (ms) |
| Utility | NBAY+3 | BOOT | Boolean | Shutter status |
| Utility | NBAY+4 | BOOT | Integer | Error number (not used !) |
| Utility | NBAY+5 | BOOT | Integer | Current optical table temp (mk) |
| Utility | NBAY+6 | BOOT | Integer | Current shield temp (mk) |
| Utility | NBAY+7 | BOOT | Integer | Current cryostat pressure (nanobar) |

Notes

- Bit significance for Timing Code Status Word (Read only)

| Bit | Significance | Comment |
| --- | --- | --- |
| 0 | Command mode | Power is on, doing idle loop execs. |
| 1 | Idle reset mode | Set if continuous reset mode active. |
| 2 | Video mode | Reserved for future video streaming. |
| 3 | Readout mode | Set if readout is in progress. |
| 4 | Timer mode | Set if elapsed timer active |
| 5 | Command waiting | Set if command waiting for service |
| 6 | Host Flag | Set if command from host |
| 7 | Command error | Set if command format error occurred |
| 8 | Even/Odd row | Used in window readout. Not relevant |

- Bit Significance for Utility Code Status Word	 (Read only)

| Bit | Significance | Comment |
| --- | --- | --- |
| 0 | Service request | 1ms Service req. pending |
| 1 | Exposure timer active | Set to request exposure timer service |
| 2 | Preflash active | Set to request preflash timer service. |
| 3 | Timeout active | Set to request command timeout service. |
| 4 | BEX in progress | Indicates exposure in progress |
| 5 | PWR in progress | Indicates PWR command in progress |