256x256 array_tests

test 10

STARE    mode readout noise (e-) =  94.338

ND_STARE mode readout noise (e-) =  96.999

DARK CURRENT (e-/sec)            =  170.090

STARE readout noise is HIGH low wrt nominal of 56e-

ND_STARE readout noise is HIGH wrt nominal of 38e-

Dark current is HIGH wrt nominal = 0-10e-/sec

test 11

STARE    mode readout noise (e-) =  89.159

ND_STARE mode readout noise (e-) =  104.417

DARK CURRENT (e-/sec)            =  160.764

test 12

STARE    mode readout noise (e-) =  99.325

ND_STARE mode readout noise (e-) =  86.507

DARK CURRENT (e-/sec)            =  159.526

64x64 array_tests

test 13

STARE    mode readout noise (e-) =  96.362

ND_STARE mode readout noise (e-) =  145.589

DARK CURRENT (e-/sec)            =  155.037

test 14

STARE    mode readout noise (e-) =  91.388

ND_STARE mode readout noise (e-) =  113.017

DARK CURRENT (e-/sec)            =  159.597

128x128 + fast readouts

test 15